## Slide 1: Radiographic Technique Formulation

## Slide 2
Purpose of Technique Charts
Standardize exposure factor selection across technologists
Promote consistent image quality regardless of the operator
Reduce repeat examinations and unnecessary patient dose
Serve as a baseline reference, not a substitute for judgment

## Slide 3
Analog
Digital
Technique Charts in Modern Radiography
Technique charts remain essential despite digital imaging
Digital systems mask exposure errors through rescaling
Poor technique still increases patient dose and noise
Technique charts help control exposure variability

## Slide 5
Relationship to ALARA
ALARA requires intentional exposure selection
Overexposure is less visually obvious in digital imaging
Technique charts reduce “exposure creep”
Dose optimization depends on consistent technique use

## Slide 6: Exposure Index (DR) - AGFA

## Slide 7: Exposure Index (DR) - AGFA

## Slide 8: Exposure Index (DR) - AGFA

## Slide 9: Exposure Index (DR) - AGFA

## Slide 10
What Technique Charts Are NOT
Not a replacement for patient assessment
Not automated decision systems
Not equivalent to APR or AEC
Not designed for every possible clinical condition

## Slide 11
What APR Is
Anatomic Programming (APR) is a computer-controlled exposure selection system built into modern radiographic generators.
Preprogrammed exposure factors based on body part and projection
Controlled by integrated microprocessor circuitry
Automatically displays suggested kVp, mA, and exposure time
Designed to improve consistency, efficiency, and workflow
Still allows manual override by the technologist
APR is essentially a digital logic system that selects exposure parameters using stored algorithms rather than human recall.

## Slide 12
What APR Is Not
APR is not autonomous and not infallible.
It does not assess patient size
It does not detect pathology
It does not replace clinical judgment
It does not guarantee correct exposure
The technologist remains fully responsible for image quality and dose optimization.

## Slide 14
Technique Chart: The Functional Equivalent
A technique chart is the conceptual and operational equivalent of APR.
Both systems serve the same core purpose:
Selecting appropriate exposure factors for a given anatomic region and projection
Technique chart → a reference system created by humans
APR → an automated exposure selection system embedded in the generator
They are functionally related, not identical

## Slide 15
APR vs AEC
APR: selects suggested technique factors
AEC: terminates exposure based on detector input
Both rely on correct technologist setup
Neither replaces critical thinking

## Slide 16
Variables Standardized in Technique Charts
A technique chart standardizes baseline exposure conditions so that image quality and patient dose remain consistent across examinations
Anatomic Part and Projection
kVp
mAs
Grid Use and Grid Ratio
Source-to-Image Receptor Distance (SID)
Image Receptor Type
Patient Thickness / Body Habitus Categories
Use of Automatic Exposure Control (AEC)?

## Slide 17
Anatomic Part and Projection
Specifies body part (e.g., chest, abdomen, extremity)
Defines projection or view (AP, PA, lateral, oblique)
Accounts for tissue thickness between projections
Why it matters: Different projections change path length and attenuation.

## Slide 18
kVp
Determines beam energy and penetration
Standardized to balance:
  Subject contrast
  Patient dose
  Scatter production

## Slide 19
mAs
Controls image receptor exposure
Standardized to ensure:
  Adequate signal
  Acceptable noise
  Consistent exposure index

## Slide 20
Grid Use and Grid Ratio
Technique charts indicate:
Whether a grid is required
Grid ratio (e.g., 8:1, 10:1, 12:1)
Why it matters:
Grids increase contrast
Require higher mAs to maintain SNR
Affect patient dose
Charts often include grid conversion factors.

## Slide 21
Source-to-Image Receptor Distance (SID)
Standardizes:
Geometric magnification
Beam intensity (inverse square law)
Consistency in exposure
Common SIDs:
40 in (100 cm)
44 in (110 cm)
72 in (180 cm) for chest

## Slide 22
Technique charts account for:
Film–screen speed (historical)
CR vs DR
Detector efficiency
Dynamic range
Why it matters:
Different receptors require different mAs A chart built for CR will overexpose DR if unchanged

## Slide 23
Patient Thickness / Body Habitus Categories
Many charts include:
Thin
Average
Large
Bariatric

## Slide 25
Bariatric Patient Category
Bariatric refers to patients whose body habitus significantly exceeds the “large” category, typically due to marked increases in adipose tissue thickness. This category is not defined by weight alone, but by radiographic thickness, tissue composition, and imaging limitations.

## Slide 26
Bariatric patients present exponential, not linear, imaging challenges. A small increase in thickness can require large increases in exposure due to the attenuation properties of adipose tissue.
Key distinguishing features:
Substantially increased part thickness (cm)
High proportion of adipose tissue
(low contrast, high scatter)
Frequent equipment and geometric limitations

## Slide 27
Scatter Radiation
Larger tissue volumes generate significantly more scatter
Degrades image contrast
Often overwhelms low-ratio grids
Implication: Higher grid ratios (10:1–16:1) or air-gap techniques may be required, depending on equipment availability.

## Slide 28
Exposure Time Constraints
Especially important for:
Motion control
Pediatric imaging
Trauma patients

## Slide 29
Technique Chart Formulation Requirements
(Digital Imaging: CR / DR)
A separate technique chart must be established for each X-ray unit, including fixed rooms and portable systems, because output and generator performance vary between units.

## Slide 30
Departmental standards for acceptable digital image quality must be defined, including:
Noise tolerance
Spatial resolution expectations
Acceptable exposure index (EI) or deviation index (DI) range
These standards prevent unnecessary repeats and dose creep.

## Slide 31
Radiation dose creep is the gradual, unintentional increase in patient radiation dose over time in digital radiography (CR/DR) because higher exposures continue to produce acceptable, or even better-looking, images without an obvious visual penalty.

## Slide 32
Why Dose Creep Happens
1. Wide Exposure Latitude
Digital detectors tolerate a broad range of exposures. Unlike film-screen, overexposure does not make the image “too dark.”
Result: Technologists don’t get immediate visual feedback that the dose is increasing.
2. Noise Bias (“Cleaner Looks Better”)
Higher exposure → higher signal-to-noise ratio (SNR). Images look smoother with less quantum mottle, so higher techniques are subconsciously favored.
3. Post-Processing Masks Errors
Automatic rescaling and processing normalize brightness and contrast, hiding overexposure.
4. Defensive Technique Selection
To avoid repeats (especially in trauma, bariatric, or uncooperative patients), technologists may “add a little extra,” which accumulates over time.

## Slide 33
What Dose Creep Is NOT
Not a single bad exposure
Not intentional overdosing
Not visible just by looking at the image
It is a systemic, workflow-level problem.

## Slide 34
Importance of Accurate Part Thickness Measurement
Accurate measurement of anatomic part thickness is essential for the effective and consistent use of exposure technique charts, particularly in digital imaging, where image appearance may mask exposure errors.
Calipers are devices used to measure part thickness and should be readily available in every radiographic room. Technique charts should clearly specify the exact location at which the measurement is to be taken to ensure consistency among technologists.
Part thickness is typically measured:
At the central ray (CR) midpoint, or
At the thickest portion of the anatomy being radiographed
Failure to measure part thickness accurately is one of the most common sources of error when using technique charts and can lead to improper detector exposure, increased noise, or unnecessary patient dose

## Slide 35
Always measure part thickness at the level of the central ray—never estimate.

## Slide 36
Types of Exposure Technique Charts
Two primary types of exposure technique charts are used in radiography:
Variable kVp / Fixed mAs
Fixed kVp / Variable mAs
Each chart type has distinct characteristics that affect measurement accuracy, radiographic contrast, patient dose, and x-ray tube heat load. The choice of chart design affects both image quality and radiation safety.

## Slide 37
Clinical Interpretation
Variable kVp / Fixed mAs
  Requires precise part thickness measurement
  Changes in kVp alter beam penetration and contrast
  Greater potential for dose variability
  Increased tube heat due to higher kVp usage
Fixed kVp / Variable mAs
  Maintains consistent radiographic contrast
  Adjusts exposure primarily through mAs
  Generally results in lower patient dose
  Reduced tube heat load
  Preferred in digital radiography for consistency and dose control
In digital imaging, fixed kVp/variable mAs charts are commonly preferred because they produce consistent image contrast while minimizing patient dose and tube heat.

## Slide 38
Variable kVp / Fixed mAs Technique Chart
A variable kVp/fixed mAs technique chart is based on the principle that beam energy (kVp) increases with anatomic part thickness, while mAs remains constant.
A baseline kVp is established for a specific anatomic part and projection.
For each 1-cm increase in part thickness, the kVp is increased by approximately 2 kVp.
The mAs is not changed, assuming consistent detector exposure.
The baseline kVp is adjusted upward or downward solely based on measured part thickness.
This technique requires accurate measurement of part thickness, as small measurement errors can significantly affect image contrast and patient dose.

## Slide 39
Determination of Baseline kVp (Digital Imaging)
In digital radiography (CR/DR), the determination of baseline kilovoltage (kVp) for each anatomic area is not standardized and may vary by department, equipment manufacturer, and detector characteristics. Unlike film-screen imaging, image appearance in digital systems is largely controlled by image processing, making correct baseline kVp selection critical for dose management rather than visual density.
The objective in establishing a baseline kVp is to select a beam energy that:
Provides adequate penetration of the anatomic part
Produces a detector exposure within the target Exposure Index (EI) range
Allows predictable adjustment using the 2-kVp per 1-cm change in tissue thickness rule

## Slide 40
In digital imaging, baseline kVp values are commonly determined experimentally using radiographic phantoms that simulate human tissue attenuation. Test exposures are evaluated based on:
Exposure Index (EI) or Deviation Index (DI)
Image noise characteristics
Penetration of the anatomy
Patient dose considerations
Once established, the baseline kVp serves as a reference value for thickness-based kVp adjustments, while maintaining consistent detector exposure and minimizing the risk of dose creep.

## Slide 41
High-frequency generators are grouped with three-phase generators for baseline kVp calculations because both produce a near-constant potential beam.

## Slide 42
Digital Imaging Context In CR/DR systems:
High-frequency generators dominate modern installations
The formula provides an initial estimate only
Final baseline kVp should be verified using:
  Phantoms
  Exposure Index (EI / DI)
  Noise evaluation
  Dose optimization

## Slide 43
Penetration depends on tissue composition, not thickness alone

## Slide 44
Wide Range of Tissue Densities
The chest contains very different tissue densities in the same image:
Air-filled lungs (very low attenuation)
Soft tissue (heart, vessels)
Bone (ribs, spine)
High kVp is required to:
Penetrate the densest structures (mediastinum and spine)
Avoid excessive contrast between ribs and lung fields
Display all anatomy within the detector’s dynamic range

## Slide 45
Long Scale (Low) Contrast Is Desired
Chest imaging requires many shades of gray, not high contrast.
High kVp:
Produces long-scale contrast
Allows visualization of subtle lung markings and pathology
Prevents ribs from overpowering lung detail
The abdomen, by contrast, benefits from shorter-scale (higher) contrast.

## Slide 46
Why the Abdomen Uses Lower kVp
The abdomen:
Is composed primarily of soft tissue and fluid
Has more uniform attenuation
Requires lower kVp to maintain soft-tissue contrast
Uses higher mAs instead of higher kVp

## Slide 47
Variable kVp / Fixed mAs Technique Chart (Digital Imaging)
Anatomic Part: Knee
Projection: AP
Image Receptor: CR
Tabletop / Bucky: Bucky
Measuring Point: Midpatella
Grid Ratio: 12:1
SID: 40 inches
Focal Spot: Small

## Slide 48
Variable kVp Technique Charts in Digital Imaging
In digital radiography (CR/DR), variable kVp technique charts may be particularly effective for pediatric patients and small extremity imaging due to the low attenuation and thin anatomy involved.
In these examinations:
Small changes in part thickness significantly affect x-ray penetration
Adjusting kVp allows precise control of beam energy while keeping mAs low
Maintaining low mAs reduces patient dose, which is critical in pediatric imaging
Digital image processing maintains consistent image appearance despite kVp variation
Because digital systems can mask exposure errors, variable kVp charts are most effective when accurate part-thickness measurements and exposure-index monitoring are used to prevent dose creep.
Although variable kVp technique charts can be effective in theory for thin anatomy and pediatric patients, they are rarely used in modern digital radiography (CR/DR).

## Slide 49
Fixed kVp / Variable mAs technique charts
These provide:
Stable contrast
Predictable detector exposure
Better dose optimization
Easier standardization

## Slide 50
Fixed kVp / Variable mAs Technique
Anatomic Part: Knee Projection: AP Image Receptor: CR Tabletop / Bucky: Bucky Measuring Point: Midpatella Grid Ratio: 12:1 SID: 40 inches Focal Spot: Small

## Slide 51
kVp remains fixed to maintain consistent radiographic contrast
mAs doubles for each 5-cm increase in part thickness
Thickness grouping reduces reliance on exact caliper precision
Designed for digital imaging, where detector exposure is controlled by mAs
Supports:
Consistent Exposure Index (EI)
Improved dose management
Departmental standardization

## Slide 54
Fixed kVp / Variable mAs Adjustment (Digital Imaging)
Once an optimal kVp has been established for a given anatomic area, fixed kVp / variable mAs technique charts compensate for changes in patient size by adjusting mAs only, while keeping kVp constant.
Because x-ray attenuation increases exponentially with tissue thickness, exposure must increase proportionally to maintain consistent detector exposure.
A commonly used guideline is:
For every 4–5 cm increase in part thickness, the mAs should be doubled.
This method provides predictable control of detector exposure, Exposure Index (EI), and patient dose in digital radiography.

## Slide 55
Variables NOT Typically Included on the Technique Chart
Pathology Not Reflected by Thickness
Patient Cooperation and Motion Risk
Casts, Splints, Dressings, and Medical Devices
Nonstandard projections
Portable and Trauma Condition
Prosthetics
Technique charts provide starting values, not final answers. The technologist must recognize clinical factors not listed on the chart and adjust exposure accordingly.