## Slide 1: AEC

## Slide 2: Digital System AEC 80 kVp
8 mAs
EI 800

## Slide 3: Digital System AEC 80 kVp
1 mAs
EI 100

## Slide 4: Digital System AEC
1 mAs
EI 100

## Slide 5: Digital System AEC
1 mAs
EI 100

## Slide 6: Digital System AEC 80 kVp
8 mAs
EI 800

## Slide 7: Digital System AEC 80 kVp
1 mAs
EI 100

## Slide 8: Digital System AEC 80 kVp
-2
4 mAs
EI 400

## Slide 9: Digital System AEC 80 kVp
-3
2 mAs
EI 200

## Slide 10: Digital System AEC 80 kVp
-5
1 mAs
EI 100

## Slide 11: Digital System AEC 80 kVp
+2
10 mAs
EI 1000

## Slide 12: Digital System AEC 80 kVp
+3
16 mAs
EI 1600

## Slide 13: Digital System AEC 80 kVp
+5
20 mAs
EI 2000

## Slide 14: Digital System AEC 80 kVp
L
32 mAs
EI 3200

## Slide 15: Digital System AEC 80 kVp
M
8 mAs
Lgm 2.27

## Slide 16: Digital System AEC 80 kVp
H
2 mAs
EI 200

## Slide 17: Digital System AEC 80 kVp
kVp increased to 95
4 mAs
EI 800

## Slide 18: Digital System AEC 80 kVp
kVp increased to 115
2 mAs
EI 800

## Slide 19: Digital System AEC 80 kVp
kVp decreased to 65
16 mAs
EI 800

## Slide 20: Digital System AEC 80 kVp
kVp decreased to 55
32 mAs
EI 800

## Slide 21: Digital System AEC 80 kVp
SID decreased to 50 cm (20”)
2 mAs
EI 800

## Slide 22: Digital System AEC 80 kVp
SID increased to 200 cm (80”)
32 mAs
EI 800

## Slide 23: Digital System AEC 80 kVp
CR off centered 8” laterally
2 mAs
EI 200

## Slide 24: Digital System AEC 80 kVp
CR off centered 8” longitudinally
7 mAs
EI 700

## Slide 25: Digital System AEC 80 kVp
Cephalic angulation 20⁰
8 mAs
EI 800

## Slide 26: Digital System AEC 80 kVp
Caudal angulation 20⁰
8 mAs
EI 800

## Slide 27: Digital System AEC 80 kVp
Vertical bucky selected
Backup 16 mAs
16 mAs
EI 1600

## Slide 28: Digital System AEC 80 kVp
Backup time 40 mAs
(TABLE BUCKY)
8 mAs
EI 800

## Slide 29: Digital System AEC 80 kVp
Backup time 4 mAs
AEC OVER!
4 mAs
EI 400

## Slide 30: Digital System AEC 80 kVp
Vertical bucky selected
Buckup time 200 mAs
200 mAs
EI 4000